AliveCor, Inc. CONFIDENTIAL

This document is electronically controlled; print copies are considered uncontrolled. This document is confidential and controlled property of AliveCor, Inc. Neither reproduction nor distribution to unauthorized persons in or outside of AliveCor, Inc. is allowed without proper authorization.

Change History

| Revision | Date | Description of Change | Author |
| --- | --- | --- | --- |
| 1 | See InfoCard | Initial Document | Shantanu Patil |
| 2 | See InfoCard | KardiaCore integration with KardiaMobile Card | Manoj Meghavaram |

Contents

1.	Background and Introduction	3

2.	Terminology	3

3.	Materials	4

3.1.	Components of System	4

3.2.	Equipment	4

4.	Sample Size	4

5.	TESTING CONFIGURATION	4

6.	Adoption Rationale	5

7.	Risk Assessment	5

8.	References	5

9.	APPENDICES	6

10.	Procedures and Acceptance Criterion (DATA SHEET)	7

10.1.	User Information	7

10.2.	Equipment Information	7

11.	Executions	8

11.1.	Execute for KardiaMobile 6L (Mandatory Execution)	8

11.2.	Execute for KardiaMobile	18

11.3          Execute for KardiaMobile Card	28

# Background and Introduction

The purpose of this Partner System Validation Protocol is for the Alivecor partner to validate that when the Kardia Core software is integrated to the partner host application, the combined software meets the software needs for the ‘KardiaMobile 6L with Kardia Core app’ System requirements (19REQ7) & KardiaCard System Requirements (21REQ1).

The Kardia Core software provides functionality for interaction with KardiaMobile 6L ECG event recorders, as well as compatibility with additional AliveCor devices, such as KardiaMobile or KardiaMobile Card,     that function via ultrasonic FM transmission or Bluetooth Low Energy mode of transmission.

This protocol includes test cases derived from the “KardiaMobile 6L with Kardia Core” System Validation Protocol (19VAP8) and KardiaCard System Validation Protocol (21VAP2) and are applicable to AliveCor ECG Event recorders that use the two above-mentioned modes of transmission and shall be executed for those devices, KardiaMobile and KardiaMobile Card in addition to KardiaMobile 6L. This Protocol applies to all compatible platforms (Android and iOS)and should be executed separately for each platform to be used by the partner.

The main scenarios that will be tested include:

- Recording a 1-Lead and a 6-Lead ECG

- Viewing ECGs on the Kardia Core App

- Viewing, printing, and emailing stored ECGs on the Kardia Core App

- Replacing the battery (Only applicable to KardiaMobile 6L)

# Terminology

| Term | Definition |
| --- | --- |
| AF | Atrial Fibrillation |
| Bradycardia | Bradycardia is a condition wherein an individual has a slow heart rate, typically defined as a heart rate of under 50 beats per minute (BPM) in adults. |
| ECG | Electrocardiogram |
| KM | KardiaMobile |
| MCP | Mobile Computing Platform |
| Must / shall | Indicates a requirement |
| NSR | Normal Sinus Rhythm |
| Should | Indicates a goal, but not a requirement |
| Host application or Host app | Any mobile software application that integrates the Kardia Core app using its API |
| Partner application or Partner app | A Host app created by a third-party that AliveCor partners with. |
| Tachycardia | Tachycardia is a condition wherein an individual has a faster heart rate than normal, typically defined as a heart rate of over 100 beats per minute (BPM) in adults. |
| Kardia Core app | Kardia Core app is an independently released and marketed entity that provides the following functions: - Manage the pairing of the Kardia hardware with the MCP - Acquire the signal using Bluetooth or ultrasonic transmission from the Kardia hardware, - Decode the ECG from the acquired data and display the ECG to the user, - Once a recording is complete, execute the KardiaAI ECG analysis and display the results of that analysis, - Provide the ability to export the ECG rhythm strip in the PDF format, so that a healthcare professional can perform a clinical review of ECG rhythm, and measure the QT and heart-rate corrected QT interval. For the purposes of this partner validation protocol, all mentions of the Kardia Core app imply execution within a partner host application which has integrated the Kardia Core app. |
| KardiaAI | An AliveCor ECG analysis software as a medical device that is distributed as a software library accessed through its API. This software can analyze ECG data from AliveCor-designed ECG event recorders. |
| KAIv2 | KardiaAI version 2 |

# Materials

NOTE: Applicable Information such as Model, SN, and Version shall be recorded in data sheets.

## Components of System

- KardiaMobile 6L Device

- KardiaMobile Device (as applicable)

- Kardia Core app integrated within a partner’s host application

- Kardia Instructions for Use (IFU)

- Kardia Core Service Manual

## Equipment

- MCP compatible with Kardia Devices

# Sample Size

Each test case will be performed on

- 1 Kardia Mobile 6L Device

- 1 KardiaMobile Device (as applicable)

- 1 KardiaMobile Card

- 1 MCP with the Kardia Core App installed

- 1 IFU (19LB65)

- 1 IFU (02LB49) (as applicable)

- 1 IFU (21LB01) (as applicable)

# TESTING CONFIGURATION

The Kardia devices, Kardia Core App, and IFU versions tested will be representative of the commercially available Kardia system.

The KardiaMobile 6L device and KardiaMobile Card will be connected via Bluetooth to the MCP with the Kardia Core App installed. KardiaMobile uses audio ultrasonic transmission to connect to the MCP with the Kardia Core App installed.

Test cases listed in Section 11 will be executed and results will be documented with a “PASS” or “FAIL” which represents whether the acceptance criteria for the test is met.   Meeting the acceptance criteria also requires that any failures and deviations found during testing are found to be acceptable, after being analyzed and dispositioned by the project team.

When applicable, test data in the form of screenshots, printed reports, and emailed reports will be saved as objective evidence.

For traceability, the requirement(s) associated with each scenario are provided in the requirements column.

# Adoption Rationale

N/A

# Risk Assessment

There are no risks identified for performing these tests.

# References

- 19DDP3 –Design and Development Plan, KardiaMobile 6L with Kardia Core App

- 19RSK7 – Risk Management Plan

- 19RSK8 – KardiaMobile 6L with Kardia Core App System Risk Assessment and Control Document

- 19RSK9 – KardiaMobile 6L with Kardia Core Risk Table

- 19REQ7 – KardiaMobile 6L with Kardia Core System Requirements

- 21VAP2 – KardiaCard System Validation Test Protocol

- IEC 62366-1:2015 – Medical devices – Part 1: Application of usability engineering to medical devices

- ISO 10993-5:2009 – Biological evaluation of Medical Devices, Part 5: Tests for cytotoxicity, in-vitro methods

- ISO 10993-10:2010 – Biological Evaluation of Medical Devices, Part 10: Tests for Irritation and Skin Sensitization

- ISO 10993-12:2012 – Biological Evaluation of Medical Devices, Part 12: Sample Preparation and Reference Materials

- FDA Guidance for Industry and Food and Drug Administration Staff (June 2016) – Use of International Standard ISO 10993-1, "Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process"

- IEC 60601-1:2005+AMD1:2012 – Medical electrical equipment – Part 1: General requirements for basic safety and essential performance, Edition 3.1

- IEC 60601-1-11:2015 – Medical electrical equipment – Part 1-11:  General requirements for basic safety and essential performance – Collateral Standard: Requirements for medical electrical equipment and medical electrical systems used in the home healthcare environment

- IEC 60601-1-2:2014: Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral Standard: Electromagnetic disturbances

- IEC 60601-2-47:2012 – Medical electrical equipment – Part 2-47: Particular requirements for the basic safety and essential performance of ambulatory electrocardiographic systems

# APPENDICES

N/A

# Procedures and Acceptance Criterion (DATA SHEET)

## User Information

| Date(s) of Test Execution |  |
| --- | --- |
| Name of Tester(s) |  |
| Initials of Tester(s) |  |
| Signature of tester |  |
| AliveCor Approval: |  |

## Equipment Information

| KardiaMobile 6L Device Hardware Version |  |
| --- | --- |
| KardiaMobile 6L Device Hardware Serial Number |  |
| KardiaMobile Device Hardware Version |  |
| KardiaMobile Device Hardware Serial Number |  |
| KardiaMobile Card Device Hardware Version |  |
| KardiaMobile Card Device Hardware Serial Number |  |
| MCP Model Number |  |
| MCP Serial Number |  |
| MCP OS Version |  |
| MCP Kardia Core App Version |  |

# Executions

## Execute for KardiaMobile 6L (Mandatory Execution)

| Case # | Requirement | Procedures | Acceptance Criteria | Result (Pass or Fail) |
| --- | --- | --- | --- | --- |
| 10 | When three electrodes are utilized to take a recording, the system shall enable users to record, store, and view (real-time and stored) a standard simultaneous 6-lead EKG (Lead I, Lead II, Lead III, aVF, aVL, and aVR) and average heart rate. | - Follow the on-screen instructions to record a 6-Lead ECG. - Save a screenshot (you may need assistance) of the real-time display of the 6-Lead ECG. - Complete the 6-Lead ECG recording. - Save a screenshot to show that the 6-Lead ECG is complete. | User is able to record a 6-Lead ECG using the Kardia Core App. |  |
|  |  |  | The real-time 6-Lead ECG displays Lead I, Lead II, Lead III, aVF, aVL, and aVR and the current heart rate. |  |
|  |  | - Log-out of the Kardia Core App and log-in again. - View and save a screenshot of the previously recorded 6-Lead ECG. | User is able to view their stored 6-Lead ECG using the software. |  |
|  |  |  | The stored 6-Lead ECG displays Lead I, Lead II, Lead III, aVF, aVL, and aVR and the average heart rate. |  |
| 20 | When two electrodes are utilized to take a recording, the system shall enable users to record, store, and view a standard 1-lead EKG and average heart rate. | - Follow the on-screen instructions to record a 1-Lead ECG. - Save a screenshot (you may need assistance) of the real-time display of the 1-Lead ECG. - Complete the 1-Lead ECG recording. - Save a screenshot of the result screen to show that the 1-Lead ECG is complete. | User is able to record a 1-Lead ECG using the Kardia Core App. |  |
|  |  |  | The real-time 1-Lead ECG is displayed with the current heart rate. |  |
|  |  | - From the result screen, choose to export the EKG as a PDF. - View and save a screenshot of the previously recorded 1-Lead ECG now in PDF form | User is able to view a secondary display of an existing 1-Lead ECG using the Kardia Core App. |  |
|  |  |  | The stored 1-Lead ECG is displayed with the average heart rate on the PDF. |  |
| 30 | System shall use BLE to enable the device to connect with Kardia Core App installed on a Mobile Computing Platform (MCP). | - Un-pair and forget any KardiaMobile 6L devices listed in available Bluetooth devices under the MCP’s Bluetooth settings. - Open the Kardia Core app and navigate to a recording screen. - Follow the on-screen instructions to record an ECG. - Save a screenshot to show that the Kardia Core App confirms that the Kardia Mobile device will be paired via Bluetooth. - Save a screenshot that the ECG is complete. | User is able to record and view their ECG using the Kardia Core App installed on the MCP when the Kardia Mobile 6L device is connected to the MCP via Bluetooth using the Kardia Core App. |  |
| 40 | System shall be able to record an EKG when device in close proximity (without being attached) to the MCP. | - Move the KardiaMobile 6L device as far away from the MCP as possible while still able to view the Kardia Core App screen. - Follow the on-screen instructions to record an ECG. - Save a screenshot that the ECG is complete. | User is able to record and view their ECG using the Kardia Core App when the KardiaMobile 6L device is in close proximity (without being attached) to the MCP. |  |
| 50 | System shall produce 30 second recordings with clinically relevant accuracy. | 1. Navigate to the recording screen for your bluetooth device 2. Confirm device is set to Single-Lead mode on recording screen 3. Begin recording an EKG in single lead mode by placing left hand fingers on the top left electrode, and right hand finger on top right electrode 4. Keep hands on EKG hardware until the countdown timer reaches zero 5. Repeat for Six lead. | The recording ends at the conclusion of the countdown, and you are taken to an EKG result screen |  |
| 60 | System shall provide assistance to user on how to obtain an ECG recording (i.e. BLE connection, how to hold device correctly, etc.). | - Create a new Kardia Core App account. - Follow the on-screen instructions to record the first ECG. - Choose the KardiaMobile 6L device. - At the recording screen observer the on screen prompts and further help that can be opened by choosing “I need help” - Save a screenshot of the video tutorial provided for how to use the KardiaMobile 6L device. - Complete the video. | Kardia Core App provides assistance to user on how to obtain a 1-Lead ECG recording including: - How to pair using Bluetooth - How to hold device in correct orientation - Correct contacts for a 1-Lead ECG - Using the countdown timer to complete the recording |  |
| 70 | System shall inform user of any conditions preventing an ECG recording or analysis (i.e. too short, interference, etc.). | 1. Navigate to the EKG recording screen 2. Begin recording a single lead EKG 3. With 10 seconds left, stop recording. 4. The app will present an alert. Choose the option to ‘Save’ | The app will present an Instant analysis screen with the result of "Too Short" |  |
| 80 | System should be compatible with MCPs that AliveCor identifies as compatible models. | 1. Start application 2. Using the touchscreen, tap the log in button 3. Using the touchscreen keyboard, enter a valid user ID and password | User is able to navigate and login using the touchscreen on the device |  |
| 90 | System shall display information about the device's battery life. | - Open the Kardia Core App. - Start an ECG recording. - Save a screenshot of the ECG recording screen that displays the KardiaMobile 6L device’s battery life. | The Kardia Core App displays the KardiaMobile 6L device’s battery life. |  |
| 180 | User shall receive the analysis determination of their recorded EKGs. | 1. complete an NSR recording at 80 bpm for 30 seconds using the Fluke patient simulator set to program #6. If simulator is not used, take a single lead reading. 2. arrive at the Instant Analysis screen. | The Instant Analysis screen should display the EKG result description in the area below the EKG |  |
| 190 | User shall be able to share the recorded EKG(s) with their physician. | - Ensure the MCP is connected to a printer. - Complete an ECG recording. - After the ECG recording completes, print the ECG. - Save the screenshot of the print confirmation screen. - Save the printed ECG. | Recorded ECGs can be printed using the Kardia Core App or MCP OS standard capabilities |  |
|  |  | - Ensure the MCP has internet connection. - Complete an ECG recording. - After the ECG recording completes, email the ECG. - Save the screenshot of the email confirmation screen. - Save the emailed ECG. | Recorded ECGs can be emailed using the Kardia Core App or MCP OS standard capabilities |  |
|  |  | - Ensure the MCP is connected to a printer. - Return to the recording result screen, choose to display a PDF, and then choose to share it to a printer - Print the ECG. - Save the screenshot of the print confirmation screen. - Save the printed ECG. | Recorded ECGs can be printed using the Kardia Core App or MCP OS standard capabilities |  |
|  |  | - Ensure the MCP has internet connection. - View a previously recorded ECG. - Email the ECG. - Save the screenshot of the email confirmation screen. - Save the emailed ECG. | Recorded ECGs can be emailed using the Kardia Core App or MCP OS standard capabilities |  |
| 230 | US accounts will not have access to KAIv2 EKG analysis results until they are under the care of a physician (only to be executed by partners implementing use of KardiaAI v2) | - Use the Kardia Core app within a host application configured to use KAIv2 for analysis - Configure the application to be identified as an implementation that has not been validated to have its users under the care of a physician. - Go through the recording flow. | The Recording flow will not initiate, and a recording cannot be taken. |  |
|  |  | - Configure the application to be considered an implementation which HAS verified with AliveCor that the users are under the care of a physician. - Begin and complete an EKG recording. | EKG will be analyzed by KAIv2 and yield KAIv2 results and identification as a v2 or Advanced determination result. |  |
| 240 | System shall provide an ECG recording and display functionality API that can be incorporated within partner host application | - Use the Kardia Core app implemented within a host application - Progress through the recording flow until arriving at the result screen | Host application will have executed the Kardia Core app using the available APIs. |  |
|  |  | - Review the contents of the Kardia Core Service manual and included appendices. | The Kardia Core Service Manual (02LB83) contains information necessary to integrate the Kardia Core app software |  |
| 250 | System shall provide software integration information including compatibility and integration methods to ensure the API can be safely and effectively incorporated. | - Review the contents of the Kardia Core Service manual and included appendices. | The Kardia Core Service Manual (02LB83) contains information to ensure the API can be safely and effectively incorporated |  |

## Execute for KardiaMobile

| Case # | Requirement | Procedures | Acceptance Criteria | Result (Pass or Fail) |
| --- | --- | --- | --- | --- |
| 20 | When two electrodes are utilized to take a recording, the system shall enable users to record, store, and view a standard 1-lead EKG and average heart rate. Note: KardiaMobile Device uses audio Ultrasonic transmission to connect to the MCP | - Follow the on-screen instructions to record a 1-Lead ECG. - Save a screenshot (you may need assistance) of the real-time display of the 1-Lead ECG. - Complete the 1-Lead ECG recording. - Save a screenshot of the result screen to show that the 1-Lead ECG is complete. | User is able to record a 1-Lead ECG using the Kardia Core App. |  |
|  |  |  | The real-time 1-Lead ECG is displayed with the current heart rate. |  |
|  |  | - From the result screen, choose to export the EKG as a PDF. - View and save a screenshot of the previously recorded 1-Lead ECG now in PDF form | User is able to view a secondary display of an existing 1-Lead ECG using the Kardia Core App. |  |
|  |  |  | The stored 1-Lead ECG is displayed with the average heart rate on the PDF. |  |
| 30 | System shall use Ultrasonic to enable the device to connect with the Kardia Core App installed on a Mobile Computing Platform (MCP). | - Open the Kardia Core app and navigate to a recording screen. - Follow the on-screen instructions to record an ECG. - Save a screenshot to show that the Kardia Core App confirms that the Kardia Mobile device is connected via ultrasonic connection. - Save a screenshot that the ECG is complete. | User is able to record and view their ECG using the Kardia Core App installed on the MCP when the Kardia Mobile device is connected to the MCP via Ultrasonic using the Kardia Core App. |  |
| 40 | System shall be able to record an EKG when the device is in close proximity (without being attached) to the MCP. | - Move the KardiaMobile device as far away from the MCP as possible while still able to view the Kardia Core App screen. - Follow the on-screen instructions to record an ECG. - Save a screenshot that the ECG is complete. | User is able to record and view their ECG using the Kardia Core App when the KardiaMobile device is in close proximity (without being attached) to the MCP. |  |
| 50 | System shall produce 30 second recordings with clinically relevant accuracy. | 1. Navigate to the recording screen for your Kardia device 2. Begin recording an EKG by placing left hand fingers on the top left electrode, and right hand finger on top right electrode 3. Keep hands on EKG hardware until the countdown timer reaches zero. | The recording ends at the conclusion of the countdown, and you are taken to an EKG result screen |  |
| 60 | System shall provide assistance to the user on how to obtain an ECG recording (i.e. connection, how to hold the device correctly, etc.). | - Create a new Kardia Core App account. - Follow the on-screen instructions to record the first ECG. - Choose the KardiaMobile device. - At the recording screen observer the on screen prompts and further help that can be opened by choosing “I need help” - Save a screenshot of the video tutorial provided for how to use the Kardia Mobile device being used. - Complete the video. | Kardia Core App provides assistance to user on how to obtain a 1-Lead ECG recording including: - How to pair connect - How to hold device in correct orientation - Correct contacts for a 1-Lead ECG - Using the countdown timer to complete the recording |  |
| 70 | System shall inform the user of any conditions preventing an ECG recording or analysis (i.e. too short, interference, etc.). | 1. Navigate to the EKG recording screen 2. Begin recording a 1-Lead EKG 3. With 10 seconds left, stop recording. 4. The app will present an alert. Choose the option to ‘Save’ | The app will present an Instant analysis screen with the result of "Too Short" |  |
| 90 | System shall display information about the device's battery life. | - Open the Kardia Core App. - Start an ECG recording. - Save a screenshot of the ECG recording screen that displays the Kardia Mobile device’s battery life. | The Kardia Core App displays the Kardia Mobile device’s battery life. |  |
| 180 | User shall receive the analysis determination of their recorded EKGs. | 1. complete an NSR recording at 80 bpm for 30 seconds using the Fluke patient simulator set to program #6. If no simulator is used, take a single lead reading. 2. arrive at the Instant Analysis screen. | The Instant Analysis screen should display the EKG result description in the area below the EKG |  |
| 190 | User shall be able to share the recorded EKG(s) with their physician. | - Ensure the MCP is connected to a printer. - Complete an ECG recording. - After the ECG recording completes, print the ECG. - Save the screenshot of the print confirmation screen. - Save the printed ECG. | Recorded ECGs can be printed using the Kardia Core App |  |
|  |  | - Ensure the MCP has internet connection. - Complete an ECG recording. - After the ECG recording completes, email the ECG. - Save the screenshot of the email confirmation screen. - Save the emailed ECG. | Recorded ECGs can be emailed using the Kardia Core App or MCP OS standard capabilities |  |
|  |  | - Ensure the MCP is connected to a printer. - Return to the recording result screen, choose to display a PDF, and then choose to share it to a printer - Print the ECG. - Save the screenshot of the print confirmation screen. - Save the printed ECG. | Recorded ECGs can be printed using the Kardia Core App or MCP OS standard capabilities |  |
|  |  | - Ensure the MCP has internet connection. - View a previously recorded ECG. - Email the ECG. - Save the screenshot of the email confirmation screen. - Save the emailed ECG. | Recorded ECGs can be emailed using the Kardia Core App or MCP OS standard capabilities |  |
| 230 | US accounts will not have access to KAIv2 EKG analysis results until they are under the care of a physician. (only to be executed by partners implementing use of KardiaAI v2) | - Use the Kardia Core app within host application configured to use KAIv2 for analysis - Configure the application to be identified as an implementation that has not been validated to have its users under the care of a physician. - Go through the recording flow. | The Recording flow will not initiate, and a recording cannot be taken. |  |
|  |  | - Configure the application to be considered an implementation which HAS verified with AliveCor that the users are under the care of a physician. - Begin and complete an EKG recording. | EKG will be analyzed by KAIv2 and yield KAIv2 results and identification as a v2 or Advanced determination result. |  |
| 240 | System shall provide an ECG recording and display functionality API that can be incorporated within partner host application | - Use the Kardia Core app implemented within a host application - Progress through the recording flow until arriving at the result screen | Host application will have executed the Kardia Core app using the available APIs. |  |
|  |  | - Review the contents of the Kardia Core Service manual and included appendices. | The Kardia Core Service Manual (02LB83) contains information necessary to integrate the Kardia Core app software |  |
| 250 | System shall provide software integration information including compatibility and integration methods to ensure the API can be safely and effectively incorporated. | - Review the contents of the Kardia Core Service manual and included appendices. | The Kardia Core Service Manual (02LB83) contains information to ensure the API can be safely and effectively incorporated |  |

## 11.3          Execute for KardiaMobile Card

| Sys Req ID / Test # | Requirement | Procedures | Acceptance Criteria | Result (Pass or Fail) |
| --- | --- | --- | --- | --- |
| 10 | When two electrodes are utilized to take a recording, the system shall enable users to record, store, and view (real-time and stored) a standard 1-lead EKG and calculated heart rate. | - Follow the on-screen instructions to record an ECG. - Save a screenshot (you may need assistance) of the real-time display of the 1-Lead ECG. | User is able to record a 1-Lead ECG using the KardiaCore App. |  |
|  |  |  | The real-time 1-Lead ECG displays an ECG and the current heart rate. |  |
|  |  | - Log-out of the KardiaCore App and log-in again. - View and save a screenshot of the previously recorded ECG. | User is able to view their stored ECG using the KardiaCore App. |  |
|  |  |  | The stored -Lead ECG displays the ECG and the average heart rate. |  |
| 20 | System shall use BLE to enable the device to connect with KardiaCore App installed on a Mobile Computing Platform (MCP). | - Un-pair and forget any KardiaCard devices listed in available Bluetooth devices under the MCP’s Bluetooth settings. - Log-in to the KardiaCore App. - Follow the on-screen instructions to record an ECG. - Save a screenshot to show that the KardiaCore App confirms that the KardiaCard device will be paired via Bluetooth. - Save a screenshot that the ECG is complete. | User is able to record and view their ECG using the KardiaCore App installed on the MCP when the KardiaCard device is connected to the MCP via Bluetooth using the KardiaCore App. |  |
| 30 | System shall be able to record an EKG when device in close proximity (without being attached) to the MCP. | - Move the KardiaCard to a distance within 2 feet of the MCP under test. - Follow the on-screen instructions to record an ECG. - Save a screenshot that the ECG is complete. | User is able to record and view their ECG using the Kardia App when the KardiaCard device is in close proximity (without being attached) to the MCP. |  |
| 50 | System shall provide assistance to user on how to obtain an ECG recording (i.e. BLE connection, etc.). | - Create a new KardiaCore App account. - Follow the on-screen instructions to record the first ECG. - Choose the KardiaCard device. - Save a screenshot of the on-screen instructions how to use the KardiaCard device. | KardiaCore App provides assistance to user on how to obtain anECG recording including: - Pairing dialogs as needed - Contact indicator - Hand placement instructions - Countdown timer to complete the recording |  |
| 250 | User shall receive the analysis determination of their recorded EKGs. | This testing is covered in the KardiaCore app for KardiaCard Software Verification Testing. | All tests pass per acceptance criteria stated in the KardiaCore app for KardiaCard Software Verification Protocol (21VEP2). |  |
| 260 | User shall be able to share the recorded EKG(s) with their physician. | - Ensure the MCP is connected to a printer. - Complete an ECG recording. - After the ECG recording completes, print the ECG. - Save the screenshot of the print confirmation screen. - Save the printed ECG. | Recorded ECGs can be printed using the KardiaCore app |  |
|  |  | - Ensure the MCP has internet connection. - Complete an ECG recording. - After the ECG recording completes, email the ECG. - Save the screenshot of the email confirmation screen. - Save the emailed ECG. | Recorded ECGs can be emailed using the KardiaCore app |  |
|  |  | - Ensure the MCP is connected to a printer. - View a previously recorded ECG. - Print the ECG. - Save the screenshot of the print confirmation screen. - Save the printed ECG. | Previously recorded ECGs can be printed using the KardiaCore app |  |
|  |  | - Ensure the MCP has internet connection. - View a previously recorded ECG. - Email the ECG. - Save the screenshot of the email confirmation screen. - Save the emailed ECG. | Previously recorded ECGs can be emailed using the KardiaCore app |  |
|  |  | This testing is covered in the KardiaCard Firmware Verification Protocol | All tests pass per acceptance criteria stated in the KardiaCard Firmware Verification Protocol (21VEP1) |  |
| 330 | For US accounts, user shall not receive KAIv2 analysis determinations of their EKGs until after they are under the care of a physician, i.e. either after a cardiologist has reviewed their EKG, or they have a connection code through KardiaPro | This testing is covered in the KardiaCore app for KardiaCard Software Verification Testing. | All tests pass per acceptance criteria stated in the KardiaCore app for KardiaCard Software Verification Protocol (21VEP2). |  |